## In Service
Instrument | Serial Number | 2nd s/n | Case | Batts. | Status | Location | Previous Deployment | Next Deployment | Depth | Notes | CD Number
Aquadopp | 5952 |  |  |  | Deployed | KEO | KEO 2017 |  | 36.5 | Deployed at KE009 11/14/11. No inductive comms. Logging internally. Recovered 7/5/12. Bulkhead connector found corroded after KEO 2015. Replaced by Nortek 12/2016. JAK | CD0001529535
Aquadopp | 6290 |  |  |  | Recovered | Lab | KEO 2016 | KEO 2018 | 16.5 | Recovered from KE008 11/14/11. Damaged during recovery. No comms. Odd comms with Flex at KE010. Sporadic serial comms, data sync issue. | CD0001687908
Aquadopp | 6808 |  |  |  | Recovered | Lab | KEO 2016 | KEO 2018 | 36.5 | Recovered from PA004 6/13/11. Broken bolt on wire clamp. Returned end cap for repair of broken bolt 10/4/11. RMA 36-10839. Repaired 11/8/11. | CD0001687901
Aquadopp | 6810 |  |  |  | Ready | Lab | Papa 2017 Spare | Papa 2018 | 35.5 | Returned from ARC 6/9/11. Unable to communicate via serial port. Tried different end cap & was able to download data. Removed inductive ferrite to replace missing half in working unit. Sent to NortekUSA for repair 10/4/11. RMA 36-10839. Returned from Nortek 11/8/11. Deployed at KE010 7/4/12. Serial comms not working, returned for repair 3/14. Damaged on PA009 by impact from instrument above. Needs repair. 9/16 | CD0001687903
Aquadopp | 8071 |  |  |  | Recovered | Lab | Papa 2016 | Papa 2018 Spare | 35.5
Aquadopp | 8473 |  |  |  | Recovered | Lab | Papa 2016 | Papa 2018 | 15.5 | IMM fw 1.14 | CD0004075716
Aquadopp | 9819 |  |  |  | Deployed | Papa | Papa 2017 |  | 35.5 | Received new 4/4/12. | CD0004054807
Aquadopp | 9980 |  |  |  | Deployed | KEO | KEO 2017 |  | 16.5 | Received new 3/5/12. Deployed at KE010 7/4/12. | CD0004054792
Aquadopp | 12026 |  |  |  | Ready | Lab | KEO 2017 Spare | KEO 2018 | 8.5 | Received new 9/19/14. Desiccant pack broke open during deployment on KEO 2015. Returned to Nortek for servic3 12/2016. JAK
Aquadopp | 12241 |  |  |  | Recovered | Lab | KEO 2016 | KEO 2018 Spare | 8.5 | Received new 9/19/14.
Aquadopp | 12690 |  |  |  | Deployed | KEO | KEO 2017 |  | 8.5 | Received new 12/18/15.
Aquadopp | 13499 |  |  |  | Deployed | Papa | Papa 2017 |  | 15.5 |  | CD0004167601
Aquadopp Profiler | 13314 |  |  |  | Deployed | Papa | Papa 2017 |  | 68 |  | CD0004225510
Aquadopp Profiler | 13317 |  |  |  | Deployed | KEO | KEO 2017 |  | 68 |  | CD0004225511
ATRH (KEO) | 51042 |  |  |  | Deployed | KEO | KEO 2017 TFLEX |  | -3 | Deployed at KE010 7/4/12.
ATRH (KEO) | 58365 |  |  |  | Recovered | Lab | KEO 2016 | KEO 2018 Spare | -3 | Did not pass RH post-cal. SK bought new RH sensor & installed 3/2012. Sent to Rotronic for testing 10/11/12.
ATRH (KEO) | 118825 |  |  |  | Needs Cal | KEO | KEO 2017 Spare | KEO 2018 | -3 | Did not pass RH post-cal. RH probe replaced. Deployed at KE010 7/4/12.
ATRH (KEO) | 133374 |  |  |  | Recovered | Lab | KEO 2016 TFLEX | KEO 2018 TFLEX | -3 | Recovered 7/5/12.
ATRH (KEO) | 133390 |  |  |  | Deployed | KEO | KEO 2017 |  | -3 | Recovered 7/5/12.
ATRH (Papa) | 91589 |  |  |  | Deployed | Papa | Papa 2017 TFLEX |  | -3
ATRH (Papa) | 104889 |  |  |  | Deployed | Papa | Papa 2017 |  | -3 | Sent to Rotronic for testing 10/11/12.
ATRH (Papa) | 112220 |  |  |  | Needs Cal | Lab | Papa 2017 Spare | Papa 2018 | -3
ATRH (Papa) | 112220 |  |  |  | Recovered | Lab | Papa 2016 TFLEX | Papa 2018 TFLEX | -3 | Sent to Rotronic for testing 10/11/12.
ATRH (Papa) | 500600 |  |  |  | Recovered | Lab | Papa 2016 | Papa 2018 Spare | -3
Buoy | PMEL 4 | [9 empty] | CD0001537473
Buoy | PMEL 5 | [9 empty] | CD0001537474
Buoy | PMEL 6 | [9 empty] | CD0001537475
Buoy | PMEL 7 | [9 empty] | CD0001537476
Camera | 33004512 | [5 empty] | KEO 2017 |  |  | Nikon CoolPix Purchased 7/15 | CD0004177725
Camera | 33008493 | [5 empty] | Papa 2017 | Papa 2018 |  | Nikon CoolPix Purchased 7/15 | CD0004177724
Camera (Video) | 13901214 | [8 empty] | Contour GPS Purchased 4/12 | CD0004054823
Computer (Laptop) | [5 empty] | JK Office |  |  |  |  | CD0004177731
Computer (ToughBook CF-30) | OCS1 |  |  |  | Ready | Lab
Computer (ToughBook CF-30) | OCS2 |  |  |  | Needs Repair | Lab |  |  |  | Screen going bad
Computer (ToughBook CF-52) | OCS3 |  |  |  | Ready | Lab
Computer (ToughBook CF-52) | OCS4 |  |  |  | Ready | Lab
Computer Monitor | [5 empty] | Lab |  |  |  | Monitor for lab computer | CD0001529494
Computer Monitor | [5 empty] | JK Office |  |  |  | JK monitor | CD0001690508
Druck Baro | 1749053 |  |  |  | Deployed | Papa | Papa 2017 TFLEX |  | -3 | Deployed at KE010 7/4/12.
Druck Baro | 2153585 |  |  |  | Recovered | Lab | KEO 2016 | KEO 2018 Spare | -3
Druck Baro | 2153676 |  |  |  | Deployed | KEO | KEO 2017 |  | -3 | Deployed at KE009 11/14/11. Recovered 7/5/12.
Druck Baro | 2153683 |  |  |  | Needs Cal | KEO | KEO 2017 Spare | KEO 2018 | -3 | Replaced broken top shield 12/13/16.
Druck Baro | 4249223 |  |  |  | Deployed | KEO | KEO 2017 TFLEX |  | -3
Druck Baro | 4252470 |  |  |  | Needs repair | Lab | KEO 2016 TFLEX | KEO 2018 TFLEX | -3
Druck Baro | 4253762 |  |  |  | Needs Cal | Lab | Papa 2017 Spare | Papa 2018 | -3
Druck Baro | 4299622 |  |  |  | Recovered | Lab | Papa 2016 TFLEX | Papa 2018 TFLEX | -3
Eppley PIR (LWR) | 38437 | PIC1001 | Plastic |  | Deployed | Papa | Papa 2017 |  | -3 | OCS
Eppley PIR (LWR) | 38438 | PIC1002 | Plastic |  | Deployed | Papa | Papa 2017 TFLEX |  | -3 | OCS
Eppley PIR (LWR) | 38439 | PIC1003 | Plastic |  | Needs Cal | Lab | Papa 2017 Spare | Papa 2018 | -3 | OCS
Eppley PIR (LWR) | 38440 | PIC1007 | Plastic |  | Needs Cal | Lab | New | Papa 2018 TFLEX |  | Purchased for TPOS NDBC Enhancement project, but bought the wrong type. Move to OCS property, and buy replacement on OCS budget.
Eppley PIR (LWR) | 38486 | PIC1004 | Plastic |  | Deployed | KEO | KEO 2017 |  | -3
Eppley PIR (LWR) | 38487 | PIC1005 | Plastic |  | Deployed | KEO | KEO 2017 TFLEX |  | -3
Eppley PIR (LWR) | 38488 | PIC1006 | Plastic |  | Needs Cal | Lab | KEO 2017 Spare | KEO 2018 | -3
Eppley PIR (LWR) | 38692 | PIC1008 |  |  | Needs Cal | Lab | New | KEO 2018 Spare
Eppley PIR (LWR) | 38693 | PIC1009 |  |  | Needs Cal | Lab | New | KEO 2018 TFLEX
Eppley PIR (LWR) | 38694 | PIC1010 |  |  | Needs Cal | Lab | New | Papa 2018 Spare
Eppley SPP (SWR) | 38427 | PIC1 | Plastic |  | Deployed | Papa | Papa 2017 |  | -3 | OCS
Eppley SPP (SWR) | 38428 | PIC2 | Plastic |  | Deployed | Papa | Papa 2017 TFLEX |  | -3 | OCS
Eppley SPP (SWR) | 38429 | PIC3 | Plastic |  | Needs Cal | Lab | Papa 2017 Spare | Papa 2018 | -3 | OCS
Eppley SPP (SWR) | 38430 | PIC4 | Plastic |  | Needs Cal | Lab | New | Papa 2018 TFLEX |  | Purchased for TPOS NDBC Enhancement project, but bought the wrong type. Move to OCS property, and buy replacement on OCS budget.
Eppley SPP (SWR) | 38431 | PIC8 |  |  | Not built | Lab | New | Papa 2018 Spare |  | Purchased for TPOS Saildrone project but not used. Transferred to OCS.
Eppley SPP (SWR) | 38432 | PIC9 |  |  | Not built | Lab | New | KEO 2018 TFLEX |  | Purchased for TPOS Saildrone project but not used. Transferred to OCS.
Eppley SPP (SWR) | 38433 | PIC10 |  |  | Not built | Lab | New | KEO 2018 Spare |  | Purchased for TPOS Saildrone project but not used. Transferred to OCS.
Eppley SPP (SWR) | 38434 |  |  |  | Not built | Lab | New |  |  | Purchased for TPOS Saildrone project but not used. Transferred to OCS.
Eppley SPP (SWR) | 38475 | PIC5 | Plastic |  | Deployed | KEO | KEO 2017 |  | -3
Eppley SPP (SWR) | 38484 | PIC6 | Plastic |  | Deployed | KEO | KEO 2017 TFLEX |  | -3
Eppley SPP (SWR) | 38485 | PIC7 | Plastic |  | Needs Cal | Lab | KEO 2017 Spare | KEO 2018 | -3
Face Plate | 01 |  |  |  | Ready | Lab | Papa 2017 Spare | Papa 2018 | -1 | Recovered from PA004 6/13/11.
Face Plate | 03 |  |  |  | Deployed | Papa | Papa 2017 |  | -1
Face Plate | 04 |  |  |  | Recovered | Lab | KEO 2016 | KEO 2018 Spare | -1 | Unused Iridium connector available for new rad sensor if needed. Deployed at KE009 11/14/11. Recovered 7/5/12.
Face Plate | 05 |  |  |  | Not Updated | Lab | KEO 2012 | Test System | -1 | Unused Iridium connector available for new rad sensor if needed.
Face Plate | 06 |  |  |  | Recovered | Lab | Papa 2016 | Papa 2018 Spare | -1
Face Plate | 07 |  |  |  | Deployed | KEO | KEO 2017 |  | -1
Face Plate | 08 |  |  |  | Ready | Lab | KEO 2017 Spare | KEO 2018 | -1 | Paros connector found corroded after PA006. Replaced 7/2013.
FLEX Box (Druck) | 0006 |  |  |  | Recovered | Lab | KEO 2016 | KEO 2018 Spare | -1 | IMM comms had 2E35 error at ARC. Reinstalled original CPU board and could not replicate problem. Tested 4/2011. Added Iridium handshaking 1/10/12. Replaced internal IMM due to intermittent 2E35s 10/17/12.
FLEX Box (Druck) | 0008 |  |  |  | Deployed | KEO | KEO 2017 |  | -1 | Deployed at KE009 11/14/11. Recovered 7/5/12. Added Iridium handshaking 10/2012.
FLEX Box (Druck) | 0009 |  |  |  | Needs Cal | Lab | KEO 2017 Spare | KEO 2018 | -1 | Returned from ARC 6/9/11. Added Iridium handhsaking 1/10/12. Deployed at KE010 7/4/12.
FLEX Box (Druck) | 0010 |  |  |  | Flakey | Lab | KEO 2010 |  |  | Showed 28 resets between 04:00 5/28/11 and 13:00 5/29/11. Multiple resets, then stopped transmitting 9/14/11. Recovered from KE008 11/14/11. Pat tested, changed firmware. Still stalls on startup without RF modem. Added Iridium handshaking 10/2012.
FLEX Box (Paros) | 0002 |  |  |  | Deployed | Papa | Papa 2017 |  | -1 | Swapped boards with 0007 after exposed to 40V. Modified boards and changed to Paros wiring. Added Iridium handshaking and hygroclip wiring. 12/22/11
FLEX Box (Paros) | 0004 |  |  |  | Recovered | Lab | Papa 2016 | Papa 2018 Spare | -1 | Recovered from PA004 6/13/11. Added Iridium handshaking. 1/10/12
FLEX Box (Paros) | 0005 |  |  |  | Ready | Lab | Papa 2017 Spare | Papa 2018 | -1 | Added Iridium handshaking 10/2012.
Gill Wind | 044001 |  |  |  | Recovered | Lab | Papa 2016 TFLEX | Papa 2018 TFLEX | -4 | Installed on new molded base. Sparton S/N N1262.
Gill Wind | 051414 |  |  |  | Needs Repair | Lab | KEO 2016 | KEO 2018 Spare | -4 | Installed on new molded base w/ Sparton comapss N701.
Gill Wind | 051415 |  |  |  | Deployed | KEO | KEO 2017 TFLEX |  | -4 | Installed on new molded base w/ Sparton comapss P7005, 9/2016.
Gill Wind | 070229 |  |  |  | Deployed | Papa | Papa 2017 TFLEX |  | -4 | Installed on new molded base w/ Sparton compass S/N K463. Compass failed. Replaced with S/N P7004 12/7/16.
Gill Wind | 073805 | N701 |  |  | Needs Cal | KEO | KEO 2017 Spare | KEO 2018 | -4 | Installed on new molded base with Sparton S/N P4738.
Gill Wind | 10510080 |  |  |  | Deployed | KEO | KEO 2017 |  | -4 | Installed on new molded base. Sparton S/N P3047.
Gill Wind | 10510081 |  |  |  | Recovered | Lab | Papa 2016 | Papa 2018 Spare | -4 | Installed on new molded base. Sparton S/N P3046.
Gill Wind | 10510082 |  |  |  | Needs Cal | Lab | Papa 2017 Spare | Papa 2018 | -4 | Installed on new molded base. Sparton S/N P3049.
Gill Wind | 11520107 |  |  |  | Deployed | Papa | Papa 2017 |  | -4 | Installed on new molded base. Sparton S/N K457.
Gill Wind | 14180060 |  |  |  | New | Lab |  |  |  | New. Not installed on molded base. No compass.
Gill Wind | 14180061 | OS-29847 |  |  | New | Lab |  | KEO 2018 TFLEX | -4
Handheld GPS1 | GPS1 |  |  |  |  | Papa
Handheld GPS2 | GPS2 |  |  |  |  | KEO |  |  |  |  | CD0004225483
HygroClip | 20044582 |  |  |  | Deployed | KEO | KEO 2017 |  | -3 | Received new 12/15
HygroClip | 20086871 |  |  |  | New | Lab |  | KEO 2018 Spare | -3 | Received new 12/16.
HygroClip | 61222482 |  |  |  | Deployed | Papa | Papa 2017 |  | -3 | Potted at Rotronic
HygroClip | 61334062 |  |  |  | Needs Cal | Lab | Papa 2017 Spare | Papa 2018 | -3 | Received as replacement for damaged sensor.
HygroClip | 61334171 |  |  |  | Recovered | Lab | Papa 2016 | Papa 2018 Spare | -3 | Received as replacement for damaged sensor.
HygroClip | 61334287 |  |  |  | Needs Cal | Lab | KEO 2017 Spare | KEO 2018 | -3 | Received as replacement for damaged sensor.
SBE37 IM TC | 6072 |  | Titanium | D | Deployed | Papa | Papa 2017 |  | 100 |  | CD0001526674
SBE37 IM TC | 6073 |  | Titanium | D | Deployed | Papa | Papa 2017 |  | 120 |  | CD0001526675
SBE37 IM TC | 6074 |  | Titanium | D | Deployed | Papa | Papa 2017 |  | 150 |  | CD0001526676
SBE37 IM TC | 6075 |  | Titanium | D | Deployed | Papa | Papa 2017 |  | 10 |  | CD0001526677
SBE37 IM TC | 6076 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 10 | Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001526678
SBE37 IM TC | 6077 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 15 | Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001526679
SBE37 IM TC | 6078 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 25 | Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001526680
SBE37 IM TC | 6079 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 35 | Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001526681
SBE37 IM TC | 6140 |  | Titanium | D | Deployed | Papa | Papa 2017 |  | 37 | Recovered from PA004 6/13/11. Returned for cal 8/31/11. Loaned to Kessler/Stalin 9/15/11. Returned damaged - to be replaced. 1/26/12 | CD0001526685
SBE37 IM TC | 6141 |  | Titanium | D | Recovered | Lab | Papa 2016 | Papa 2018 | 25 | Recovered from PA004 6/13/11. Returned for cal 8/31/11. Loaned to Kessler/Stalin 9/15/11. Returned damaged - to be repaired. 1/26/12 | CD0001526686
SBE37 IM TC | 6142 |  | Titanium | D | Recovered | Lab | Papa 2016 | Papa 2018 | 30 | Recovered from PA004 6/13/11. Calibrated 11/28/11. | CD0001526687
SBE37 IM TC | 6145 |  | Titanium | D | Recovered | Lab | Papa 2016 | Papa 2018 | 37 | Recovered from PA004 6/13/11. Calibrated 11/28/11. | CD0001526690
SBE37 IM TC | 6146 |  | Titanium | D | Recovered | Lab | Papa 2016 | Papa 2018 | 45 | Recovered from PA004 6/13/11. Calibrated 11/28/11. | CD0001526691
SBE37 IM TC | 7786 |  | Titanium | D | Recovered | Lab | Papa 2016 | Papa 2018 | 60 | Recovered from PA004 6/13/11. Calibrated 11/28/11. | CD0001690225
SBE37 IM TC | 7787 |  | Titanium | D | Recovered | Lab | Papa 2016 | Papa 2018 | 80 | Recovered from PA004 6/13/11. Calibrated 11/28/11. | CD0001690226
SBE37 IM TC | 7788 |  | Titanium | D | Deployed | Papa | Papa 2017 |  | 200 | Recovered from PA004 6/13/11. Calibrated 11/28/11. | CD0001690227
SBE37 IM TC | 7789 |  | Titanium | D | Recovered | Lab | Papa 2016 | Papa 2018 | 100 | Recovered from PA004 6/13/11. Calibrated 11/28/11. | CD0001690228
SBE37 IM TC | 7790 |  | Titanium | D | Recovered | Lab | Papa 2016 | Papa 2018 | 120 | Recovered from PA004 6/13/11. Calibrated 11/28/11. | CD0001690229
SBE37 IM TC | 7791 |  | Titanium | D | Recovered | Lab | Papa 2016 | Papa 2018 | 150 | Recovered from PA004 6/13/11. Calibrated 11/28/11. | CD0001690230
SBE37 IM TC | 7792 |  | Titanium | D | Recovered | Lab | Papa 2016 | Papa 2018 | 200 | Recovered from PA004 6/13/11. Calibrated 11/28/11. | CD0001690231
SBE37 IM TC | 7793 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 10 | Recovered from KE008 11/14/11. Deployed at KE010 7/4/12. | CD0001690232
SBE37 IM TC | 8419 |  | Plastic | D | Deployed | Papa | Papa 2017 |  | 25 | Received new 3/29/11. Plastic housing good to 250m. | CD0001690559
SBE37 IM TC | 8420 |  | Plastic | D | Deployed | Papa | Papa 2017 |  | 30 | Received new 3/29/11. Plastic housing good to 250m. | CD0001690570
SBE37 IM TC | 8421 |  | Plastic | D | Recovered | Lab | Papa 2016 | Papa 2018 Spare | 10 | Received new 3/29/11. Plastic housing good to 250m. | CD0001690571
SBE37 IM TC | 8422 |  | Plastic | D | Deployed | Papa | Papa 2017 |  | 45 | Received new 3/29/11. Plastic housing good to 250m. | CD0001690572
SBE37 IM TC | 8423 |  | Plastic | D | Deployed | Papa | Papa 2017 |  | 60 | Received new 3/29/11. Plastic housing good to 250m. | CD0001690588
SBE37 IM TC | 8424 |  | Plastic | D | Deployed | Papa | Papa 2017 |  | 80 | Received new 3/29/11. Plastic housing good to 250m. | CD0001690573
SBE37 IM TC | 9412 |  | Plastic | AA | Recovered | Lab | Papa 2016 | Papa 2018 | 20 | Received new 3/16/12. Plastic housing good to 250m. Missing cell guard. Case came open during recovery from PA006. 6/28/2013
SBE37 IM TC | 12229 |  | Plastic | AA | Recovered | Lab | Papa 2016 | Papa 2018 | 14 | New 6/14. | CD0001467326
SBE37 IM TC | 12517 |  | Plastic | AA | Ready | Lab | Papa 2017 Spare | Papa 2018 | 10 | Received new 9/2014
SBE37 IM TC | 13248 |  | Plastic | AA | Deployed | Papa | Papa 2017 |  | 14 | New 4/15. | CD0004177697
SBE37 IM TC | 20087 |  | Plastic | AA | New | OCS Lab
SBE37 IM TC | 20433 |  | Plastic | AA | OCS Lab - | OCS Lab
SBE37 IM TCP | 7092 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 Spare | 525 | Calibrated 2/15/11. Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001524722
SBE37 IM TCP | 7093 |  | Titanium | AA | Deployed | KEO | KEO 2017 |  | 75 | AA battery pack reinstalled 3/2015. | CD0001524723
SBE37 IM TCP | 7094 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 100 | Calibrated 2/16/11. Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001524724
SBE37 IM TCP | 7095 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 125 | Calibrated 2/16/11. Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001524725
SBE37 IM TCP | 7096 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 150 | Calibrated 2/16/11. Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001524726
SBE37 IM TCP | 7097 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 175 | Calibrated 2/16/11. Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001524727
SBE37 IM TCP | 7098 |  | Titanium | AA | Deployed | KEO | KEO 2017 |  | 225 | AA battery pack reinstalled 3/2015. | CD0001524728
SBE37 IM TCP | 7099 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 275 | Calibrated 2/16/11. Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001524729
SBE37 IM TCP | 7100 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 325 | Calibrated 2/17/11. Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001524730
SBE37 IM TCP | 7101 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 525 | Calibrated 2/17/11. Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001524731
SBE37 IM TCP | 7102 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 15 | Recovered from KE008 11/14/11. | CD0001524732
SBE37 IM TCP | 7103 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 25 | Recovered from KE008 11/14/11. Deployed at KE010 7/4/12. | CD0001524733
SBE37 IM TCP | 7104 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 35 | Recovered from KE008 11/14/11. Missing conductivity cell shield & inlet tube. Deployed at KE010 7/4/12. | CD0001524734
SBE37 IM TCP | 7105 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 50 | Recovered from KE008 11/14/11. Deployed at KE010 7/4/12. | CD0001524735
SBE37 IM TCP | 7106 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 75 | Recovered from KE008 11/14/11. Missing conductivity cell & inlet tube. Deployed at KE010 7/4/12. | CD0001524736
SBE37 IM TCP | 7107 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 100 | Recovered from KE008 11/14/11. Deployed at KE010 7/4/12. | CD0001524737
SBE37 IM TCP | 7108 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 125 | Calibrated 2/17/11. Deployed at KE010 7/4/12. | CD0001524738
SBE37 IM TCP | 7780 |  | Titanium | D | Deployed | KEO | KEO 2017 |  | 425 | Logging not started on deployment 9/11. No data. Recovered from KE008 11/14/11. Deployed at KE010 7/4/12. | CD0001690182
SBE37 IM TCP | 7781 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 175 | Recovered from KE008 11/14/11. Deployed at KE010 7/4/12. | CD0001690183
SBE37 IM TCP | 7782 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 225 | Started reporting constant data after 22:00 4/27/11. Recovered from KE008 11/14/11. Deployed at KE010 7/4/12. | CD0001690184
SBE37 IM TCP | 7783 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 275 | Started reporting constant data after 20:00 5/26/11. Recovered from KE008 11/14/11. Deployed at KE010 7/4/12. | CD0001690185
SBE37 IM TCP | 7784 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 | 325 | Started reporting constant data after 12:00 3/13/11. Data output turned to 1E+35 after 13:00 3/29/11. Recovered from KE008 11/14/11. Unable to establish comms. Deployed at KE010 7/4/12. | CD0001690186
SBE37 IM TCP | 7785 |  | Titanium | D | Ready | Lab | KEO 2017 Spare | KEO 2018 | 525 | Started reporting constant data after 21:00 4/19/11. Recovered from KE008 11/14/11. Deployed at KE010 7/4/12. | CD0001690187
SBE37 IM TCP | 9413 |  | Plastic | AA | Recovered | Lab | KEO 2016 | KEO 2018 | 150 | Rceived new 3/16/12. Plastic housing good to 250m. Deployed at KE010 7/4/12.
SBE37 IM TCP | 12519 |  | Plastic | AA | Deployed | KEO | KEO 2017 |  | 50 | Received new 9/2014
SBE37 IM TCP | 16293 |  | Plastic | AA | New | Lab | New - Received 10/2017 |  | 425 | Depth rating 350m | CD0004323606
SBE37 SM TCP | 10503 |  | Titanium | AA | Deployed | Papa | Papa 2017 |  | 4250 | Owned by OceanSITES.
SBE37 SM TCP | 10504 |  | Titanium | AA | Deployed | KEO | KEO 2017 |  | 5700 | Owned by OceanSITES.
SBE37 SM TCP | 11926 |  | Titanium | AA | Recovered | Lab | KEO 2016 |  | 5700 | New 4/2014. OceanSITES instrument for deep ocean measurement.
SBE37 SM TCP | 12509 |  | Titanium | AA | Recovered | Lab | Papa 2016 | Papa 2018 | 4250 | New 9/2014 OceanSITES instrument for deep ocean measurement.
SBE37 SMP | 3802 |  | Titanium | D | Deployed | KEO | KEO 2017 TFLEX |  | 1 | Upgraded 7/2010. Deployed at KE010 7/4/12.
SBE37 SMP | 4512 |  | Titanium | AA | Needs Updating | Lab | Papa 2013 |  | 1 | Unable to upgrade. FW is v2, which will allow pump to run out of water.
SBE37 SMP | 4562 |  | Titanium | D | Recovered | Lab | KEO 2016 | KEO 2018 Spare | 1 | Recovered from KE008 11/14/11. | CD0001526684
SBE37 SMP | 7088 |  | Plastic | D | Needs Repair | Lab | KEO 2017 Spare | KEO 2018 | 1 | Calibrated 4/15/11. Deployed at KE009 11/14/11. Recovered 7/5/12. Problem noted with bad memory blocks. Memory re-mapped at Seabird. | CD0001529598
SBE37 SMP | 7089 |  | Plastic | D | Deployed | Papa | Papa 2017 TFLEX |  | 1 | Calibrated 2/16/11. Recovered 6/3/12. Problem noted with bad memory blocks. Memory re-mapped at Seabird. | CD0001529599
SBE37 SMP | 7090 |  | Plastic | D | Deployed | Papa | Papa 2017 |  | 1 | Recovered from PA004 6/13/11. Calibrated 11/28/11. | CD0001529600
SBE37 SMP | 11552 |  | Plastic | AA | Recovered | Lab | KEO 2016 TFLEX | KEO 2018 TFLEX | 1 | New 2/2014
SBE37 SMP | 11553 |  | Plastic | AA | Recovered | Lab | Papa 2016 TFLEX | Papa 2018 TFLEX | 1 | New 2/2014
SBE37 SMP | 11554 |  | Plastic | AA | Deployed | KEO | KEO 2017 |  | 1 | New 2/2014
SBE37 SMP | 11555 |  | Plastic | AA | Needs Repair | Lab | Papa 2017 Spare | Papa 2018 | 1 | New 2/2014
SBE37 SMP | 12520 |  | Plastic | AA | Recovered | Lab | Papa 2016 | Papa 2018 Spare | 1 | Received new 9/2014
SBE39 IM T | 3283 |  | Plastic |  | Loaned to EDD | EDD | KEO 2014 |  |  | Recovered from PA004 6/13/11. Calibrated 11/28/11. Loaned to Scott for TELOS testing 2/2018
SBE39 IM T | 3285 |  | Plastic |  | Recovered | Lab | KEO 2016 | KEO 2018 | 20 | Deployed at KE009 11/14/11. Recovered 7/5/12.
SBE39 IM T | 3286 |  | Plastic |  | Needs batts. | Lab | KEO 2016 Spare | Questionable reliability - See Notes | 40 | Calibrated 2/17/11. Deployed at KE009 11/14/11. Recovered 7/5/12. Sent to KEO 2015 but not deployed. Dead battery. Sent to KEO 2016 but not deployed. Dead battery.
SBE39 IM T | 3287 |  | Plastic |  | Needs Cal | Lab | KEO 2015 |  | 40 | Battery failed mid-deployment on KE013. Memory pointer reset to 0, and had to request SBE download the data. They said they fixed the problem. Loaned to DZ 3/30/17 -- Returned
SBE39 IM TP | 4358 |  | Titanium |  | Deployed | KEO | KEO 2017 |  | 20 | Recovered from KE008 11/14/11. Temp probe bent. Deployed at KE010 7/4/12.
SBE39 IM TP | 4359 |  | Titanium |  | Deployed | KEO | KEO 2017 |  | 40 | Calibrated 11/28/11. Deployed at KE010 7/4/12.
SBE39 IM TP | 4360 |  | Titanium |  | Recovered | Lab | KEO 2016 | KEO 2018 | 375 | Recovered from KE008 11/14/11. Temp probe bent. Deployed at KE010 7/4/12.
SBE39 IM TP | 4361 |  | Titanium |  | Ready | Lab | KEO 2017 Spare | KEO 2018 | 5 | Returned from ARC 6/9/11. Calibrated 11/28/11. Deployed at KE010 7/4/12.
SBE39 IM TP | 4377 |  | Titanium |  | Deployed | KEO | KEO 2017 |  | 5 | Returned from ARC 6/9/11. Calibrated 11/28/11. | CD0001526692
SBE39 IM TP | 4378 |  | Titanium |  | Recovered | Lab | KEO 2016 | KEO 2018 | 475 | Recovered from KE008 11/14/11. | CD0001526693
SBE39 IM TP | 4379 |  | Titanium |  | Deployed | KEO | KEO 2017 |  | 375 |  | CD0001526694
SBE39 IM TP | 4380 |  | Titanium |  | Deployed | Papa | Papa 2017 |  | 5 |  | CD0001526696
SBE39 IM TP | 4857 |  | Titanium |  | Recovered | Lab | KEO 2016 | KEO 2018 | 40 | Calibrated 2/17/11.
SBE39 IM TP | 4858 |  | Titanium |  | Needs Cal | Lab | Papa 2014 |  | 300 | Loaned to TMBA 4/16 - Returned 3/17 Deploy on ADCP mooring
SBE39 IM TP | 4859 |  | Titanium |  | Deployed | KEO | KEO 2017 |  | 475 | Calibrated 2/17/11. Deployed at KE009 11/14/11. Recovered 7/5/12.
SBE39 IM TP | 4860 |  | Titanium |  | Recovered | Lab | Papa 2016 | Papa 2018 Spare | 5 | Calibrated 2/17/11. Deployed at KE009 11/14/11. Recovered 7/5/12.
SBE39 IM TP | 4861 |  | Titanium |  | Recovered | Lab | KEO 2016 | KEO 2018 Spare | 5
SBE39 IM TP | 4862 |  | Titanium |  | Ready | Lab | Papa 2017 Spare | Papa 2018 | 5 | Recovered from PA004 6/13/11. Calibrated 11/28/11.
SBE39 IM TP | 4863 |  | Titanium |  | Deployed | Papa | Papa 2017 |  | 175 | Recovered from PA004 6/13/11. Calibrated 11/28/11.
SBE39 IM TP | 4864 |  | Titanium |  | Deployed | Papa | Papa 2017 |  | 300 | Recovered from PA004 6/13/11. Calibrated 11/28/11.
SBE39 IM TP | 4865 |  | Titanium |  | Recovered | Lab | Papa 2016 | Papa 2018 | 175 | Recovered from KE008 11/14/11.
SBE39 IM TP | 4866 |  | Titanium |  | Recovered | Lab | Papa 2016 | Papa 2018 | 300 | Recovered from KE008 11/14/11.
TFLEX Tube | 2001 |  |  |  | Deployed | Papa | Papa 2017 TFLEX |  | -1 | New 4/2014 Evidence of water intrusion when recovered from PA008. Salt crystals beneath RF antenna.
TFLEX Tube | 2002 |  |  |  | Deployed | KEO | KEO 2017 TFLEX |  | -1 | New 4/2014
TFLEX Tube | 2003 |  |  |  | Needs Cal | Lab | Papa 2017 Spare | Papa 2018 TFLEX | -1 | New 4/2014
TFLEX Tube | 2004 |  |  |  | Needs Cal | Lab | KEO 2017 Spare | KEO 2018 TFLEX | -1 | New 4/2014
TFLEX Tube | 2005 |  |  |  | Recovered | Lab | Papa 2016 TFLEX | Papa 2018 Spare | -1 | New 4/2014
TFLEX Tube | 2006 |  |  |  | Recovered | Lab | KEO 2016 TFLEX | KEO 2018 Spare | -1 | New 4/2014
Workhorse Sentinel ADCP | 14605 |  |  |  | Deployed | Papa | Papa 2017 |  | 2 | Updated fw to 50.41 10/2016 Returned to RDI for basic service 10/2016. See service documents in OCS lab for details.
Workhorse Sentinel ADCP | 14607 |  |  |  | Recovered | Lab | Papa 2016 | Papa 2018 | 2 | Updated fw to 50.41 5/10/16

## EDD-TELOS
Instrument | Serial Number | Case | Batts. | Status | Location | Previous Deployment | Next Deployment | Depth | Notes | CD Number
Gill Wind | 14180058 |  |  | New | Lab |  | TELOS Testing |  | Delivered to EDD for TELOS testing. 1/4/18 OS4000-T Compass S/N 000029787
HygroClip | 20086880 |  |  | New | Lab |  | TELOS Testing | -3 | Delivered to EDD for TELOS testing. 1/4/18

## Retired
Instrument | Serial Number |  |  | Status | Location | Last Deployment | Notes
ATLAS Tube (KEO) | 684 |  |  |  | Returned to TAO | KEO 2012 | Deployed at KE010 7/4/12.
ATLAS Tube (KEO) | 688 |  |  |  | Returned to TAO | KEO 2013
ATLAS Tube (KEO) | 756 |  |  |  | Returned to TAO | KEO 2013 Spare | Recovered 7/5/12.
ATLAS Tube (Papa) | 685 |  |  | Needs Repair | Returned to TAO | PAPA 2010 | Not working. Possible damage due to battery leak.
ATLAS Tube (Papa) | 689 |  |  |  | Returned to TAO | Papa 2012 | Use spare battery from 2011.
ATLAS Tube (Papa) | 731 |  |  |  | Returned to TAO | Papa 2013 | Modified to make Papa tube. 12/22/11
ATLAS Tube (Papa) | 757 |  |  |  | Returned to TAO | Papa 2013 Spare
ATRH (Papa) | 118814 |  |  |  | Returned to TAO | Papa 2014 | AT cal 7/19/11. RH cal 8/1/11. Needs pre RH cal before deployment.
ATRH (Papa) | 118816 |  |  |  | Returned to TAO | Papa 2015 Spare | AT cal 3/28/11. RH cal 3/25/11. Stale AT cal. Use as spare. Needs pre RH cal before deployment.
AquaDopp | 5954 |  |  | Broken/Retired | Lab | KEO 2014 | Returned for repair 2/7/11 on RMA 36-10805. Received from Nortek 3/11/11 with new end connector. Deployed at KE009 11/14/11. Recovered 7/5/12. Cracked and flooded on KE-012. Completely destroyed internally. | CD0001529536
AquaDopp | 6809 |  |  | Flooded | Excessed | Papa 2015 | Recovered from KE008 11/14/11. End cap broken off during recovery. Broken inductive clamp. Repaired under RMA 36-10889 2/12. Broken bolt and ferrite while trying to mount on Papa wire 6/12. End cap found broken off when recovered from PA009, 7/5/16. | CD0001687902
Computer (Desktop) |  |  |  |  | Excessed |  | CNSD Ticket 7962 completed 5/1/17 | CD0001700827
Druck Baro | 2153648 |  |  | Broken/Retired | Lab
DVS | 00012 |  |  | Broken/Retired | 32 Storage |  | Needs upgrade
DVS | 00014 |  |  | Broken/Retired | 32 Storage |  | Needs upgrade. Loaned to RC at UW APL 2/2/11. Returned to OCS 5/19/11.
DVS | 00015 |  |  | Broken/Retired | 32 Storage |  | Needs upgrade
DVS | 10094 |  |  | Returned | TRDI | PAPA 2011 | Loaner. Returned 2/4/11 on RMA 73133. Received on new loan agreement 4/27/11 with new S/N 15635. Returned to TRDI 7/9/2012.
DVS | 10097 |  |  | Broken/Retired | 32 Storage | KEO 2010 | No comms after 0000 6/16/11. Recovered from KE008 11/14/11.
DVS | 10098 |  |  | Broken/Retired | 32 Storage |  | Returned from upgrade with new S/N 15074. Was to be deployed at KE009, but unable to establish serial or inductive comms.
DVS | 10716 |  |  | Broken/Retired | TRDI | KEO 2009 | Loaner. Returned 2/4/11 on RMA 73134.
DVS | 8922 |  |  | Broken/Retired | TRDI | PAPA 2010 | Loaner (formerly 10610) Recovered from PA004 6/13/11. Returned to TRDI 7/1/11.
Face Plate | 02 |  |  | Do Not Deploy | Lab | KEO 2010 | Recovered from KE008 11/14/11. Corrosion on contacts due to water in well.
Face Plate | 09 |  |  | Do Not Deploy | Pat's Office
Face Plate | 10 |  |  | Do Not Deploy | Lab
FLEX Box (Druck) | 0011 |  |  | Needs Repair | Pat's Office
FLEX Box (Hygroclip) | 0007 |  |  | Broken/Retired | Lab |  | Swapped boards into 0002. Contains non-working boards exposed to 40V. Loaned to EDD in 2014 for DART/PICO comparison test.
Gill ATLAS | 42214 |  |  | Broken/Retired | Unknown |  | Water damage
Gill ATLAS | 62174 |  |  | Broken/Retired | Lost | PAPA 2010 | Lost on recovery from PA004 6/13/11.
Gill Wind 2 | 044002 |  |  | Broken/Retired | Lost | PAPA 2010 | Lost on recovery from PA004 6/13/11.
Gill Wind 6 | 08170013 |  |  | Broken/Retired | Lab |  | Flooded
Gill Wind | 073804 |  |  | Retired |  | KEO 2013 | Flooded on KE011. Retired.
Gill Wind | 042213 |  |  | Lost | KEO | KEO 2015 TFLEX | Installed on new compass base. Sparton S/N K899. 8/1/14 Lost during recovery of KE013.
Gill Wind | 08170010 |  |  | Broken/Retired | KEO | KEO 2016 | Flooded. Wind & compass damaged beyond repair.
HygroClip | 60835859 |  |  | Retired | Rotronic | KEO 2013 | Replaced under warranty
HygroClip | 60835896 |  |  | Retired | Rotronic | Papa 2013 | Replaced under warranty
HygroClip | 60947413 |  |  | Retired | Rotronic | KEO 2013 | Replaced under warranty
HygroClip | 60947644 |  |  | Retired | Rotronic | Papa 2013 | Replaced under warranty
HygroClip | 60947424 |  |  | Retired | Rotronic |  | Replaced under warranty
HygroClip | 60947661 |  |  | Retired | Rotronic |  | Replaced under warranty
HygroClip | 61222400 |  |  | Broken/Retired | Garbage | KEO 2014 | Potted at Rotronic for 2014 deployment. RH failed at KE012.
HygroClip | 61365505 |  |  | Lost |  | KEO 2016 |  | -3 | Received as replacement for damaged sensor.
Load Cell | 52451 |  |  | Broken/Retired | Excessed | PAPA 2011 | Load Cell Central, drifted Pins bent
Load Cell | 52452 |  |  | Broken/Retired | Excessed |  | Load Cell Central, drifts
Load Cell | 52453 |  |  | Broken/Retired | Excessed | KEO 2011 | Load Cell Central, stopped working. Recovered 7/5/12. Bad cable? Likely flooded. Values don't change. 9/3/12.
Load Cell | A0608173 |  |  | Broken/Retired | EDD | KEO 2012 | 3PS Pancake, potted Deployed at KE010 7/4/12. Failed. Given to EDD for analysis.
Load Cell | A0608174 |  |  | Broken/Retired | Excessed
Load Cell | A0608175 |  |  | Broken/Retired | EDD | Papa 2012 | 3PS Pancake, repaired & potted EDD requested for forensics 6/2013.
Load Cell | A0907049 |  |  | Broken/Retired | Excessed | KEO 2010 | 3PS Pancake, stopped working after 3 days
SBE37 IM TCP | 7091 | Titanium | D | Lost |  | KEO 2016 |  | 425 | Calibrated 2/15/11. Deployed at KE009 11/14/11. Recovered 7/5/12. | CD0001524721
SBE37 IM TC | 12550 |  | Plastic | Deployed | Papa | Papa 2017 | Damaged beyond repair. Replaced under warranty. | 20 | Received new 9/2014
SBE51 IM TC | 0004 |  | Plastic | Retired | Bldg. 32 | Papa 2013 | SBE can no longer recover data.
SBE51 IM TC | 0005 |  | Plastic | Needs Cal | Lab | Papa 2013 | SBE can no longer recover data.
Vaisala WXT520 | B2140048 |  |  | Recovered | KEO | KEO 2014 | Sparton compass N701. 3/25/10 Deployed by JAMSTEC crew on repair cruise - 2/19/2011. Worked sporadically. Recovered from KE008 11/14/11. Wind cal done 1/5/12. No problems found. Does not seem to need repair. Test thoroughly. Deployed at KE010 7/4/12.
Vaisala WXT520 | B2170004 |  |  | Broken/Retired | Vaisala |  | New sensor installed on #2 Sparton compass base. Compass S/N K463. Repeatedly failed wind cal. Contacted Vaisala Technical Support 5/17/11. Returned to Vaisala for replacement 9/20/11.
Vaisala WXT520 | C4710007 |  |  | Broken/Retired | Vaisala |  | Wind/rain/compass OK. No BP/AT/RH. Returned 2/4/11 Vaisala Service Request 28764. Returned from repair and installed on new compass base. Sparton S/N P3048. 3/7/11 Winds stopped working. Returned for eval, SRN 32560. 4/8/11. Vaisala tech found no problem. Received at PMEL 5/4/11. Reinstalled on compass base. Failed wind cal 5/17/11. Contacted Vaisala Technical Support. Returned to Vaisala for replacement 9/20/11.
Vaisala WXT520 | D3840025 |  |  | Retired | Lab | KEO 2013 | Not functioning after storm on KE011.
Vaisala WXT520 | D3840026 |  |  | Retired | Lab | Papa 2013 | Sparton compass K899. 3/8/10 Recovered from PA004 6/13/11. Cracked & not working. Compass removed 1/11/13.
Vaisala WXT520 | D3840028 |  |  | Recovered | Lab | Papa 2014 | Flooded. No data output. C100 compass not working. Returned 2/4/11 Vaisala Service Request 28764. Returned from repair and installed on new compass base. Sparton S/N P3049. 3/7/11 Deployed at PA005. 6/11/11 Compass removed 1/11/13
Vaisala WXT520 | G4020102 |  |  | Needs Repair | Lab | KEO 2011 | Received new 10/7/11. Sparton compass K463 10/10/11. Deployed at KE009 11/14/11. Recovered 7/5/12. Compass removed. AT, RH, BP channel not working 2/21/13.
Vaisala WXT520 | G4020103 |  |  | LOST | Lost | Papa 2012 | Received new 10/7/11. Sparton compass P3048 10/10/11. Lost on PA006. Cable sheared, mast and micarta plate gone. 6/18/2013

## TPOS-NDBC
Instrument | Serial Number | Case | Batts. | Status | Location | Notes | CD Number
Aquadopp | 13128 |  |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Marked as PMEL property and shipped to Regina Moore 7/16. | CD0004167464
Aquadopp | 13138 |  |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Marked as PMEL property. Shipped to Regina Moore 11/16. | CD0004167465
Aquadopp | 13149 |  |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Marked as PMEL property. Shipped to Regina Moore 11/16. | CD0004167466
Aquadopp | 13159 |  |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Marked as PMEL property. Shipped to Regina Moore 11/16. | CD0004167467
Aquadopp | 13687 |  |  |  |  | Shipped to Regina 5/8/17 | CD0004323557
Aquadopp | 13690 |  |  |  |  | Shipped to Regina 5/8/17 | CD0004323558
Aquadopp | 13736 |  |  |  |  | Shipped to Regina 5/8/17 | CD0004323559
Aquadopp | 13750 |  |  |  |  | Shipped to Regina 5/8/17 | CD0004323560
Aquadopp | 13757 |  |  |  |  | Shipped to Regina 5/8/17 | CD0004323561
Aquadopp | 13793 |  |  |  |  | Shipped to Regina 5/8/17 | CD0004323562
Aquadopp Profiler | 12202 |  |  |  |  | Received 5/1/17, shipped to Regina 5/4/17 | CD0004225506
Aquadopp Profiler | 12215 |  |  |  |  | Received 5/1/17, shipped to Regina 5/4/17 | CD0004225507
Aquadopp Profiler | 12278 |  |  |  |  | Received 5/1/17, shipped to Regina 5/4/17 | CD0004225508
Aquadopp Profiler | 12347 |  |  |  | NDBC | Received 5/1/17 - Returned to Nortek Reordered and received again 2/27/18 Shipped with external battery case to Regina 3/6/18 | CD0004323659
Aquadopp Profiler | 12630 |  |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Marked as PMEL property and shipped to Regina Moore 7/16. | CD0004167463
Aquadopp Profiler | 13099 |  |  | New | NDBC | Sent as loaner, then kept as initial delivery of next order. Never at PMEL, but KG sent pictures on 2/28/17. | CD0004225509
Aquadopp Profiler | 13895 |  |  |  |  | Received 7/23/18 | CD0004338734
Aquadopp Profiler | 13903 |  |  |  |  | Received 7/23/18 | CD0004338735
Eppley PIR (LWR) | 38449 | Metal |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 12/16.
Eppley PIR (LWR) | 38450 | Metal |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 12/16.
Eppley PIR (LWR) | 38451 | Metal |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 12/16.
Eppley PIR (LWR) | 38452 | Metal |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 12/16.
Eppley SPP (SWR) | 38483 | Metal |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 12/16.
Eppley SPP (SWR) | 38489 | Metal |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 12/16.
Eppley SPP (SWR) | 38490 | Metal |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 12/16.
Eppley SPP (SWR) | 38491 | Metal |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 12/16.
Rain Gauge | 2361 |  |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Marked as PMEL property and shipped to Regina Moore 6/16.
Rain Gauge | 2362 |  |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Marked as PMEL property and shipped to Regina Moore 6/16.
Rain Gauge | 2363 |  |  | New | NDBC | Purchased for TPOS NDBC Enhancement project. Marked as PMEL property and shipped to Regina Moore 6/16.
SBE37 IM TC | 15369 | Plastic | AA | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017. | CD0004225467
SBE37 IM TC | 15382 | Plastic | AA | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017. | CD0004225468
SBE37 IM TC | 15383 | Plastic | AA | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017. | CD0004225469
SBE37 IM TC | 15384 | Plastic | AA | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017. | CD0004225470
SBE37 IM TC | 15385 | Plastic | AA | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017. | CD0004225471
SBE37 IM TC | 15396 | Plastic | AA | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017. | CD0004225472
SBE37 IM TC | 15400 | Plastic | AA | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017. | CD0004225473
SBE39 IM T | 8214 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM T | 8215 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM T | 8216 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM T | 8217 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM T | 8221 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM T | 8224 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM T | 8225 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM T | 8226 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM T | 8227 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM T | 8228 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM T | 8229 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM T | 8230 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM TP | 8283 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.
SBE39 IM TP | 8284 | Plastic |  | New | PMEL | Purchased for TPOS NDBC Enhancement project. Shipped to Regina Moore 3/2017.

## TPOS-Saildrone
Instrument | Serial Number | Status | Location | Previous Deployment | Next Deployment | Purchased For | CD Number | Notes | Factory Cal Date | PMEL Cal Date
Eppley SPP (SWR) | 38431 |  |  | New |  | TPOS - Saildrone |  | Moved to OCS inventory | 2016-07-21
Eppley SPP (SWR) | 38432 |  |  | New |  | TPOS - Saildrone |  | Moved to OCS inventory | 2016-07-21
Eppley SPP (SWR) | 38433 |  |  | New |  | TPOS - Saildrone |  | Moved to OCS inventory | 2016-07-21
Eppley SPP (SWR) | 38434 |  |  | New |  | TPOS - Saildrone |  | Moved to OCS inventory | 2016-07-22
Eppley PIR (LWR) | 38441 | Modified by EDD | OCS | TPOS-1 |  | TPOS - Saildrone |  |  | 2016-07-29 | 2017-07-21
Eppley PIR (LWR) | 38442 | Modified by EDD | OCS | TPOS-1 |  | TPOS - Saildrone |  |  | 2016-07-29 | 2017-07-31
Eppley PIR (LWR) | 38443 | Modified by EDD | Saildrone | New | TPOS-2 | TPOS - Saildrone |  |  | 2016-07-29 | 2018-08-03
Eppley PIR (LWR) | 38444 | Modified by EDD | Saildrone | New | TPOS-2 | TPOS - Saildrone |  |  | 2016-07-29 | 2018-08-03
DynaMax SPN1 | A1593 | Masked | Saildrone | TPOS-1 | TPOS-2 | TPOS - Saildrone | CD0004167506 |  | 2016-07-21
DynaMax SPN1 | A1594 | Shadow Mask removed | Saildrone | TPOS-1 | TPOS-2 | TPOS - Saildrone | CD0004167507 |  | 2016-07-21
DynaMax SPN1 | A1595 | Shadow Mask removed | Saildrone | TPOS-1 | TPOS-2 | TPOS - Saildrone | CD0004167508 |  | 2016-07-21
DynaMax SPN1 | A1596 | Ready | Lab | Not used |  | TPOS - Saildrone | CD0004167509 | Large spread in detector outputs. Returned to manufacturer for analysis. Received at PMEL 11/2017. | 2017-11-02
DynaMax SPN1 | A1275 | Borrowed from ITAE | Saildrone | TPOS-1 | TPOS-2 | ITAE |  | Borrowed from ITAE / FOCI | 2014-09-22
DynaMax SPN1 | A1971 | New | Lab |  |  | TPOS - Saildrone |  | Digital S/N reads 1913
DynaMax SPN1 | A1972 | New | Lab |  |  | TPOS - Saildrone |  | Digital S/N reads 1914
DynaMax SPN1 | A1973 | New | Lab |  |  | TPOS - Saildrone |  | Digital S/N reads 1915
DynaMax SPN1 | A1974 | New | Lab |  |  | TPOS - Saildrone |  | Digital S/N reads 1916
SBE56 | 8821 | New | Saildrone | New |  | TPOS - Saildrone |  |  | 2018-08-02
SBE56 | 2142 | Borrowed from CO2 |  | CO2 glider |  | CO2
SBE56 | 2143 | Borrowed from CO2 |  | CO2 glider |  | CO2

## EPIC
OCS Inventory Based on EPIC Purchases
August, 2011
| Starting EPIC Inventory (1/2004) | IO Losses | Remaining EPIC Inventory (3/2006) | KEO Losses | JKEO Losses | Papa Losses | Additional Other Losses | OCS Purchases | Total OCS Inventory (8/2011) | Current OCS Needs
Tube | 5 | 2 | 3 | 0 | 0 | 0 | 0 | 2 | 5 | 6
RMY Wind | 4 | 2 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0
ATRH | 5 | 2 | 3 | 0 | 0 | 0 | 0 | 0 | 3 | 9
SWR | 13 | 2 | 11 | 0 | 0 | 0 | 0 | 0 | 11 | 9
LWR | 13 | 1 | 12 | 0 | 0 | 0 | 0 | 0 | 12 | 9
Rain | 7 | 2 | 5 | 0 | 0 | 0 | 0 | 0 | 5 | 5
Paros BP | 11 | 1 | 10 | 0 | 0 | 0 | 3 | 0 | 7 | 3
SSC | 15 | 2 | 13 | 0 | 0 | 0 | 0 | 0 | 13 | 5
T | 40 | 7 | 33 | 2 | 0 | 0 | 0 | 0 | 31 | 0
TC | 68 | 11 | 57 | 13 | 2 | 10 | 0 | 0 | 32 | 0
TCV | 22 | 6 | 16 | 3 | 1 | 0 | 1 | 0 | 11 | 0
TP | 10 | 4 | 6 | 6 | 0 | 0 | 0 | 0 | 0 | 0
SONTEK | 25 | 5 | 20 | 3 | 0 | 2 | 0 | 0 | 15 | 0
Tower | 5 | 2 | 3 | 0 | 0 | 0 | 0 | 0 | 3 | 0
Toroid | 5 | 2 | 3 | 0 | 0 | 0 | 0 | 0 | 3 | 0
Bridle | 5 | 2 | 3 | 0 | 0 | 0 | 0 | 0 | 3 | 0
Nylon (reels) | 30 | 12 | 18 | 10 | 0 | 10 | 0 | Unknown | -2 | 22
Note on E6: Jennifer Keene: KE007 (1)
Note on H11: Jennifer Keene: From HPF e-mail 1/16/2007
Note on E13: Jennifer Keene: KE003 (2)
Note on E14: Jennifer Keene: KE003 (12), KE005 (1)
Note on F14: Jennifer Keene: KE004 (2)
Note on G14: Jennifer Keene: PA001 (1), PA002 (9)
Note on E15: Jennifer Keene: KE003 (3)
Note on F15: Jennifer Keene: KE004 (1)
Note on H15: Jennifer Keene: 2N 140W Rec 2/28/2005
Note on E16: Jennifer Keene: KE003 (5). KE006 (1)
Note on E17: Jennifer Keene: KE003 (3)
Note on G17: Jennifer Keene: PA001 (1), PA002 (1)
Note on E21: Jennifer Keene: KE002 (5), KE003 (5)
Note on G21: Jennifer Keene: PA002 (5), PA003 (5)

## EPIC2
OCS Inventory Based on EPIC Purchases
August, 2011
| Starting EPIC Inventory (1/2004) | OCS Half of EPIC | KEO Losses | JKEO Losses | Papa Losses | Additional Other Losses | OCS Purchases | Total OCS Inventory (8/2011) | Current OCS Needs
Tube | 5 | 3 | 0 | 0 | 0 | 0 | 2 | 5 | 6
RMY Wind | 4 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0
ATRH | 5 | 3 | 0 | 0 | 0 | 0 | 0 | 3 | 9
SWR | 13 | 7 | 0 | 0 | 0 | 0 | 0 | 7 | 9
LWR | 13 | 7 | 0 | 0 | 0 | 0 | 0 | 7 | 9
Rain | 7 | 4 | 0 | 0 | 0 | 0 | 0 | 4 | 5
Paros BP | 11 | 6 | 0 | 0 | 0 | 3 | 0 | 3 | 3
SSC | 15 | 8 | 0 | 0 | 0 | 0 | 0 | 8 | 5
T | 40 | 20 | 2 | 0 | 0 | 0 | 0 | 18 | 0
TC | 68 | 34 | 13 | 2 | 10 | 0 | 0 | 9 | 0
TCV | 22 | 11 | 3 | 1 | 0 | 1 | 0 | 6 | 0
TP | 10 | 5 | 6 | 0 | 0 | 0 | 0 | -1 | 0
SONTEK | 25 | 13 | 3 | 0 | 2 | 0 | 0 | 8 | 0
Tower | 5 | 3 | 0 | 0 | 0 | 0 | 0 | 3 | 0
Toroid | 5 | 3 | 0 | 0 | 0 | 0 | 0 | 3 | 0
Bridle | 5 | 3 | 0 | 0 | 0 | 0 | 0 | 3 | 0
Nylon (reels) | 30 | 15 | 10 | 0 | 10 | 0 | Unknown | -5 | 22
Note on D6: Jennifer Keene: KE007 (1)
Note on G11: Jennifer Keene: From HPF e-mail 1/16/2007
Note on D13: Jennifer Keene: KE003 (2)
Note on D14: Jennifer Keene: KE003 (12), KE005 (1)
Note on E14: Jennifer Keene: KE004 (2)
Note on F14: Jennifer Keene: PA001 (1), PA002 (9)
Note on D15: Jennifer Keene: KE003 (3)
Note on E15: Jennifer Keene: KE004 (1)
Note on G15: Jennifer Keene: 2N 140W Rec 2/28/2005
Note on D16: Jennifer Keene: KE003 (5). KE006 (1)
Note on D17: Jennifer Keene: KE003 (3)
Note on F17: Jennifer Keene: PA001 (1), PA002 (1)
Note on D21: Jennifer Keene: KE002 (5), KE003 (5)
Note on F21: Jennifer Keene: PA002 (5), PA003 (5)

## Sheet1
Status | Location
Deployed | PMEL
Recovered | OCS Lab
OCS Lab | PAPA
OCS Lab - Calibrated | KEO
OCS Lab - Needs Cal | SEABIRD
In for Cal/Repair
Testing
New